GENERAL

  - Section includes:

    - Online sensor to detect mineral oils in water.

  - Measurement Procedures

    - The method of detection is by UV fluorescent measurement for polycyclic aromatic hydrocarbons (PAH)

      - Excitation wavelength: 254 nm

      - Emission (measurement) wavelength: 360 nm

  - Alternates

    - Other methods of such as those that require oil to be extracted into an organic solvent prior to measurement are not acceptable.

    - Other methods of such as those that require addition of a detergent surfactant to a sample prior to measurement are not acceptable.

    - Other methods of such as those that require measure turbidity of the sample (surface scattering) are not acceptable.

  - System Description

    - Performance Requirements

      - Range

        - Low: 0 to 50 ppb and 0 to 500 ppb in relation to PAH calibration standard, corresponding to 0.1 to 1.5 ppm and 0.1 to 15 ppm of oil calibration standard

        - High: 0 to 500 ppb and 0 to 5000 ppb in relation to PAH calibration standard, corresponding to 0.1 to 15 ppm and 0.1 to 150 ppm of oil calibration standard

      - Resolution: 0.1 ppb (μg/L) PAH in the lowest measuring range

      - Limit of detection: 1 ppb PAH

      - Reproducibility: 2.5 % of measurement value at a constant temperature

      - Accuracy: 5 % of measurement value ± 2 % from measurement range limit at a constant temperature

      - Response time: 10 seconds (T90)

  - Certifications

    - CE certified

  - Environmental Requirements

    - Operational Criteria

      - Temperature:

        - Sample: 1 to 40 °C (34 to 104 °F)

        - Ambient: -5 to 45 °C (23 to 113 °F), 
Sensor wetted by at least half from the measuring medium: -25 to 55 °C (-13 to 131 °F)

      - Pressure:

        - Sensor: 30 bar maximum

        - Flow cell: 1 bar maximum

        - Ball valve insertion mounting hardware: 6 bar maximum

      - Sensor distance, wall to ground: 100 mm (3.94 in.) minimum

  - Warranty

    - The product includes a one-year warranty from the date of shipment.

    - Optional service agreement: annual preventive maintenance (Hach Service Partnership Program)

  - Maintenance Service

    - Inspection interval: every 2 years

    - Lamp replacement interval: every 4 years

    - Routine maintenance: clean the optics as conditions require

PRODUCTS

  - Brand

    - HACH

      - Model FP 360 sc Oil-in-Water Online Sensor

  - Manufactured Unit

    - Display units: ppb, ppm, μg/L, mg/L

    - Calibration: Factory calibrated with UV fluorescence standard or process calibration with results of a grab sample analysis. The sensor compensates for the interference effects of ambient light and UV lamp output fluctuations.

    - The sensor operates reagent-free without sample conditioning in the range from 0 to 200 ppm total suspended solids in the sample.

    - The sensor can be connected to a HACH sc controller.

    - Connection information:

      - Sensor: 8-pin (PUR)

      - Controller: M12

  - Equipment

    - Materials:

      - Measurement sensor:

        - Housing: stainless steel (1.4571 / AISI 316Ti) or titanium

        - Optic bracket: polyoxymethylene (POM)

        - Housing bolts: stainless steel (1.4571 / AISI 316Ti)

        - Measurement window: synthetic quartz glass (Suprasil)

        - Gaskets:

          - Housing: Viton®

          - Measurement window: nitrile butadiene rubber (NBR)

      - Flow cell:

        - Housing: polyoxymethylene (POM)

        - Installation board: polyvinyl chloride (PVC)

        - Gaskets: nitrile butadiene rubber (NBR)

        - Fittings: nickel-plated brass

      - Hanging ring (Shackle): Stainless steel (1.4301 / AISI 304)

      - Chain mounting hardware: Stainless steel (1.4571 / AISI 316Ti )

      - Ball valve insertion mounting hardware: components in contact with the sample are made from stainless steel (1.4408 / CF8M or 1.4571 / 316Ti), PTFE and NBR. 
Flange sets are either made from stainless steel (1.4571 / AISI 316Ti ) or carbon steel (1.0037)

  - Components

    - Standard equipment:

      - Sensor with cable and hanging ring (shackle)

    - Required equipment:

      - HACH sc controller

    - Optional equipment:

      - Flow cell or chain mounting hardware or ball valve insertion mounting hardware

      - HACH high output air blast cleaning unit (HOAB) (only in combination with chain mounting hardware)

    - Dimensions:

      - Sensor:

        - Without plugs or suspension pin: 68 x 306 mm (2.68 x 12.05 in.)

        - With suspension pin: 68 x 399 mm (2.68 x 15.71 in.)

        - With optional cleaning unit: 68 x 413 mm (2.68 x 16.26 in.)

      - Flow cell:

        - Without fittings: 98 x 98 x 150 mm (3.86 x 3.86 x 5.91 in.)

        - Installation board: 600 x 300 x 10 mm (23.62 x 11.81 x 0.39 in.)

      - Cable: 1.5 or 10 m (4.9 or 32.8 ft.)

    - Weight:

      - Sensor:

        - Stainless steel: 2.8 kg (6.1 lbs.)

        - Titanium: 1.8 kg (4.0 lbs.)

      - Flow cell:

        - 0.6 kg (1.3 lbs.) approximate

        - With installation board: 2.0 kg (4.4 lbs.) approximate

      - Chain mounting hardware:

        - 5 kg (11 lbs.) approximate

      - Ball valve insertion mounting hardware:

        - 18 kg (40 lbs.) approximate

  - Accessories

    - Extension cable (up to total maximum length of 40 m (131 ft.))

EXECUTION

  - Preparation

    - The sensor may need to be installed with additional accessories depending on its application.

      - Open channels, shafts, tanks (suspended solid matter 200 mg/L maximum):

        - Install with chain mount kit

        - Sensor with or without cleaning unit

      - Measurement media with low suspended solid matter and continuous, low sample throughput:

        - Install with flow cell

        - Sensor without cleaning unit

      - Closed pipe loops, tank walls (suspended solid matter 200 mg/L maximum)

        - Install with ball valve insertion mounting hardware

        - Sensor without cleaning unit

  - Installation

    - Contractor will install the sensor in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician upon request (additional costs).

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION